## Slide 1: FETAL HEART RATE MONITOR
Varun Rao
Partners: Grace Lee, Matt Brown
Client: Professor Naegle

## Slide 2: Project Scope
Distinguish between multiple (at least 2) fetal heart rates and maternal heart rate
Minimize constant care from healthcare provider
Ensure comfort for mother

## Slide 3: Design specifications
Specification | Description
Implementation | Our device should be able to isolate and differentiate between at least 2 fetal heart rates and the maternal heart rate. Additionally, both fetal heart rates and maternal heart rate should be outputted at a frequency of 100 Hz.
Accuracy | Fetal heart rate and maternal heart rate should be accurate to within 10%
Comfort | No repositioning required, setup time < 5 minutes and removal time < 1 minute for health practitioner
Preprocessing | Output heart rate must be displayed in real time, which is dependent on minimizing amount of signal processing used and signal-noise ratio
Cost | Less than existing solutions (Barnes-Jewish uses a $5,000 machine)
Lightweight | Recording equipment should be < 3lbs
Safety | Meet standards IEC 60601 – 1 (Safety and Effectiveness of Medial Equipment)

## Slide 4: Potential Design Solutions
Electrocardiography
  Pans-Tompkins with Array of Recording Electrodes
  Multivariate Singular Spectrum Analysis (MSSA)
  Fast Independent Component Analysis
  Sensor Selection with Array of Recording Electrodes
Phonocardiography
  Abdominal Microphone
  Abdominal Microphone + Open Air Microphone
  Array of Abdominal Microphone + Open Air Microphone
Doppler Ultrasound
  Doppler Ultrasound Transducer
  Transducer with Moving Instructions

## Slide 5: Potential Design Solutions
Electrocardiography
  Pans-Tompkins with Array of Recording Electrodes
  Multivariate Singular Spectrum Analysis (MSSA)
  Fast Independent Component Analysis
  Sensor Selection with Array of Recording Electrodes
Phonocardiography
  Abdominal Microphone
  Abdominal Microphone + Open Air Microphone
  Array of Abdominal Microphone + Open Air Microphone
Doppler Ultrasound
  Doppler Ultrasound Transducer
  Transducer with Moving Instructions

## Slide 6: Pans-tompkins
Filtering to remove baseline wandering
Pans-Tompkins Algorithm
  Derivative to detect based on slope
  Squaring to emphasize QRS complex
  Integration to smooth out multiple peaks per complex
  Thresholding to detect peaks
Isolate maternal ECG
Repeat Pans-Tompkins for fetal ECG

## Slide 7: Multivariate Singular Spectrum Analysis (MSSA)
Decomposition
  Embedding – transform to higher dimensional space
  Singular Value Decomposition – decompose into multiple orthogonal components
Reconstruction
  Grouping – group components using eigentriple grouping
  Reconstruction – diagonal averaging to obtain Hankel matrix which can then be converted into a time series
Noisy ECG (top), Isolated ECG (middle), Isolated Noise (bottom)

## Slide 8: Fast Independent Component Analysis (fica)
Preprocessing
  Remove baseline wandering and power line interference at 50 Hz (if necessary)
fICA algorithm
  Assume input signal is a linear combination of weighted input signals
  Determine demixing matrix to obtain original signals
Postprocessing
  Filter out high frequency noise

## Slide 9: Sensor Selection
Utilizes array of recording electrodes
Best recording electrode output chosen based on minimizing cost function
  Cost function measures differences between each electrode and a reference electrode
  Reference electrode output is textbook ECG signal
  High costs would imply high noise or output with elements of other ECG signals

## Slide 10: Potential Design Solutions
Electrocardiography
  Pans-Tompkins with Array of Recording Electrodes
  Multivariate Singular Spectrum Analysis (MSSA)
  Fast Independent Component Analysis
  Sensor Selection with Array of Recording Electrodes
Phonocardiography
  Abdominal Microphone
  Abdominal Microphone + Open Air Microphone
  Array of Abdominal Microphone + Open Air Microphone
Doppler Ultrasound
  Doppler Ultrasound Transducer
  Transducer with Moving Instructions

## Slide 11: PHonocardiography
Fast fourier transform (FFT) applied to input signal
Envelope Generation
  Squaring FFT to demodulate the signal
  Filtering out high frequency content
  Down sampling to reduce aliasing
  Square root to reverse scaling distortion
Threshold applied to envelope to obtain periodicity and rate of fetal heart signal
FHR signal in real time (TL). Average magnitude spectrum (TR). Matched filter coefficient (BL). Inverse FFT (BR)

## Slide 12: Abdominal Condenser Microphone
Condenser microphone placed near abdomen
Collects audio signal of fetal heart
Utilizes the piezoelectric effect to convert mechanical pressure of sound waves to electrical signal
Signal processed to detect fetal heart rate

## Slide 13: Abdominal Microphone with Open Air Microphone
Abdominal microphone records acoustic signal from fetal heart
Open air microphone used to detect air noise signal
Noise signal weighted as shown below
d(k) = s(k) + n(k)
y(k) = n’(k) * w(k)
e(k) = d(k) - y(k)
w(k+1) = w(k) + e(k)*n’(k)
Noise is removed from input signal

## Slide 14: Potential Design Solutions
Electrocardiography
  Pans-Tompkins with Array of Recording Electrodes
  Multivariate Singular Spectrum Analysis (MSSA)
  Fast Independent Component Analysis
  Sensor Selection with Array of Recording Electrodes
Phonocardiography
  Abdominal Microphone
  Abdominal Microphone + Open Air Microphone
  Array of Abdominal Microphone + Open Air Microphone
Doppler Ultrasound
  Doppler Ultrasound Transducer
  Transducer with Moving Instructions

## Slide 15: Doppler Ultrasound
Quadrature Demodulation
  Accounts for the phase error between sending and receiving signals
  Sender and receiver must share a clock
  Crosstalk, where the I and Q signals bleed into each other, needs to be avoided
Autocorrelation
  Dot product of signal with itself delayed by a time interval

## Slide 16: Array of Doppler Ultrasound Transducers
Minimize repositioning
One transducer is bound to get the best signal
Signal strength used to choose the best transducer
Autocorrelation to determine the periodicity of the signal
Thresholding to determine heart rate

## Slide 17: Positional Feedback Doppler Transducer
Transducer will indicate repositioning is required using display
As transducer is moved, display will indicate one of the cardinal directions where signal strength is high
When fetal heart rate can be outputted, display will indicate that no further repositioning is required

## Slide 18: Pugh Analysis
Specification | Weight | Pans-Tompkins | Fast ICA | Sensor Selection Algorithm | Abdominal Array and Open Air | Multiple Doppler Ultrasound
Implementation | 10 | 7 | 10 | 8 | 9 | 9
Accuracy | 8 | 7 | 8 | 5 | 8 | 6
Comfort | 8 | 9 | 4 | 9 | 9 | 2
Preprocessing | 9 | 9 | 8 | 6 | 8 | 6
Cost | 4 | 10 | 10 | 10 | 7 | 4
Lightweight | 6 | 8 | 8 | 8 | 5 | 3
Safety | 8 | 10 | 10 | 10 | 8 | 8
Total |  | 447 | 436 | 414 | 420 | 292

## Slide 19: Pan-Tompkins with Array of Recording Electrodes
Utilizes multiple recordings taken from abdominal patch of electrodes
Sensor selection can be based on signal strength or utilizing our designed sensor selection algorithm
Detection of maternal QRS complex using Pans-Tompkins and removing it by averaging Q and S values
Detecting QRS complex and utilizing R-R peak interval to determine fetal heart rate

## Slide 20: Proposed Budget
All of our materials can be obtained from the BME labs.
Cost shown below only required if we were to upgrade our design materials.
Array of electrodes - $100
Abdominal patch - $20
Physical fetal model - $20
Total - $140

## Slide 21: Questions